CLAMP MX240/2040

________________________________________________________________________________

Introduction:

This instrument driver provides programming support for CLAMP MX240/2040.

It contains functions for opening, configuring, taking measurements from, and

closing the instrument.

________________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

For RS-232 instrument drivers:

-  the instrument is connected to the RS-232 interface.

-  the COM port, baud rate, parity, and timeout supplied to the

initialize function must match the settings of the instrument.

________________________________________________________________________________

Error and Status Information:

Error codes and warnings are returned as the return value of each instrument

driver function. A program should examine this value after each call to an

instrument driver function to determine if any error occurred.  The general

meaning of the return value is as follows:

Value                   Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

Possible error codes and their meanings are listed with the corresponding

instrument driver function.

________________________________________________________________________________

How To Use This Document:

This document is intended to be used as a programming reference manual.

It describes each function in the

CLAMP MX240/2040

instrument. The functions are listed in alphabetical order with a

description of the function, C syntax of the function, a description of each

parameter, and possible error codes.

________________________________________________________________________________

Function Tree Layout:

CLAMP MX240/2040                       Function Name:

Initialize                                  mx240_init

Configure

Voltage                                  mx240_voltage

Current                                  mx240_current

Power                                    mx240_power

Energy                                   mx240_energy

Resistance                               mx240_resistance

Change mode

Timer and Cosinus                        mx240_timecos

Talker                                   mx240_talker

DataLogStore                             mx240_datalogstore

Surveillance                             mx240_surveillance

Hold                                     mx240_mem

LCD                                      mx240_lcd

Measurement and control

Configuration                            mx240_status

Parameters                               mx240_param

Measurement                              mx240_data

Close                                       mx240_close

________________________________________________________________________________

CLAMP MX240/2040

This instrument driver provides programming

support for clamps (MX240, MX2040).

Functions and Classes:

1. Initialize - Open the RS232 COM port,

initializes the instrument and stores the

configuration.

2. Configure - Configure the measurements for:

2.1.  Voltage

2.2.  Current

2.3.  Power

2.4.  Energy

2.5.  Resistance

3. Change mode - Configure the following modes:

3.1. Timer & Cosinus Reference

3.2. Talker mode

3.3. DataLog Store

3.4. Surveillance mode

3.5. Hold mode

3.6. LCD

4. Measure - Take a measurement from the

instrument in the following mode:

4.1 MIN

4.2 MAX

4.3 AVG

4.4 PEAK

4.5 First Display

4.6 Second Display

4.7 One address of DataLog

4.8 All addresses of DataLog

5. Close - close the RS232 COM port.

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

mx240_close

int mx240_close (void);

Purpose

Takes the instrument off-line.

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

221   Unable to close instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

300   Function not available

301   Wrong position of the rotary switch

302   Two arguments are incompatible

303   Measurement is overloaded (OL)

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_current

int mx240_current (int _1stFunction, int range, int _2ndFunction);

Purpose

Switches the instrument to current measurement.

Parameter List

_1stFunction

Variable Type       int

Specifies the type of measurement on first display.

Variable Type: integer

Valid Range:   0 -  AC+DC

1 -  AC

2 -  DC

3 -  CF

Default Value: 0 -  AC+DC

range

Variable Type       int

Specifies the range of the measurement

Variable Type:  integer

Variable Range:  8 - Auto Range

0 - 20 A or 200 A

1 - 200 A or 2000 A

Default Value:   8 - Auto Range

_2ndFunction

Variable Type       int

Specifies the type of measurement on second display.

Variable Type: integer

Valid Range:   0 -  None

1 -  Hz

4 -  V AC+DC

5 -  V AC

6 -  V DC

8 -  A AC+DC

Default Value: 0 -  None

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

-1   First Function out of range

-2   Range out of range

-3   Second Function out of range

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

300   Function not available

301   Wrong position of the rotary switch

302   Two arguments are incompatible

303   Measurement is overloaded (OL)

304   Incompatible with SURV mode

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_data

int mx240_data (int measure, int dataLogAddress, double *result, int *type);

Purpose

Define the measurements to execute :

MIN,MAX,AVG,PEAK,First Display,Second Display,DataLog.

MIN : minimum value stored in SURVeillance mode.

MAX : maximum value stored in SURVeillance mode.

AVG : sliding value stored in SURVeillance mode.

PEAK: peak value stored in SURVeillance mode.

1st Display: last value measured on first display.

2nd Display: last value measured on second display.

One Datalog: One value stored in DataLog.

All Datalog: All values stored in Datalog.

Parameter List

measure

Variable Type       int

Specifies the type of the measurement.

Variable Type:  integer

Variable Range:  0 - MIN

1 - MAX

2 - AVG

3 - PEAK

4 - First Display

5 - Second Display

6 - One Datalog Address

7 - All Datalog Addresses

Default Value:   4 - First Display

Notes :

MIN : minimum value stored in SURVeillance mode.

MAX : maximum value stored in SURVeillance mode.

AVG : sliding value stored in SURVeillance mode.

PEAK: peak value stored in SURVeillance mode.

All Datalog Addresses are stored in a variable :

char mx240_buff[525][25];

dataLogAddress

Variable Type       int

Specifies the address of Datalog.

Variable Type:   int

Variable Range: 1 to 512

Default Value:   1

result

Variable Type       double (passed by reference)

Displays the results of the measurement in

Scientific Format. If you want to display the

results in Floating Point Format, Use the menu

Options/Edit Panel then double click on the

Measurement control and set the Display Format

Floating Point. Use the menu Options/Operate Panel

to take control.

The unities are :

VOLT        for voltage

AMPERE      for current

OHM         for resistance

Hz          for frequency

F           for capacity

WATT,var,VA for power

Wh          for energy

Variable Type:  Double

Note: If Error 300 is displayed (Not available),

this result is not valid.

type

Variable Type       int (passed by reference)

Displays the configuration of the measurement.

1 => Hz    10 => A AC+DC   20 => V AC+DC   34 => VA

3 => F     11 => A AC      21 => V AC      44 => Watt

8 => cosL  12 => A DC      22 => V DC      54 => var

9 => cosC  13 => CF A      23 => CF V      64 => Wh

74 => Ohm

Variable Type: integer

Note: If Error 300 is displayed (Not available),

this result is not valid.

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

-1   Type of measurement out of range

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

300   Function not available

301   Wrong position of the rotary switch

302   Two arguments are incompatible

303   Measurement is overloaded (OL) or not ready

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_datalogstore

int mx240_datalogstore (int dataLogStore);

Purpose

Specifies the DATALOG mode.

Parameter List

dataLogStore

Variable Type       int

Specifies the DataLog mode to store the current value.

Variable Type:  integer

Variable Range:  0 - Off

1 - On

Default Value:   0

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

-1   Timer value out of range

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_energy

int mx240_energy (int _2ndFunction, int range);

Purpose

Switches the instrument to energy measurement.

Parameter List

_2ndFunction

Variable Type       int

Specifies the type of measurement on second display.

Variable Type: integer

Valid Range:   0 -  None

4 -  V AC+DC

8 -  A AC+DC

Default Value: 0 -  None

range

Variable Type       int

Specifies the range of the measurement

Variable Type:  integer

Variable Range:  8 - Auto Range

0 - 2.000k or 20.00k

1 - 20.00k or 200.0k

2 - 200.0k or 2000 k

Default Value:   8 - Auto Range

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

-2   Range out of range

-3   Second Function out of range

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

300   Function not available

301   Wrong position of the rotary switch

302   Two arguments are incompatible

303   Measurement is overloaded (OL)

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_init

int mx240_init (int COMPort);

Purpose

Initializes the instrument module by opening a

COM port, and reads the status of the CLAMP.

Parameter List

COMPort

Variable Type       int

Specifies the COM port for the RS232.

Variable Type: integer

Valid Range:   1 to 8 COM1 to COM8

Default Value: 1 - COM1

Return Value

Displays the results of the instrument driver

function call:

Variable Type: integer

Error Codes:

0   Success

-1   COM port out of range

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

239   Error in configuring timeout

240   Instrument timed out

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_lcd

int mx240_lcd (int status);

Purpose

Changes the status of the LCD. It can display a

measurement or all the segments.

Parameter List

status

Variable Type       int

Specifies the status of the display:

- Normal display: displays a measurement

- Full display  : all segments are displayed

Variable Type:  integer

Variable Range:  0 - Normal display

1 - full display

Default Value:   0 - Normal display

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

-1   Status out of range

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_mem

int mx240_mem (int hold);

Purpose

Specifies the HOLD mode.

Parameter List

hold

Variable Type       int

Specifies the HOLD mode to fixe the display to the

current value.

Variable Type:  integer

Variable Range:  0 - Off

1 - On

Default Value:   0

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

-1   Timer value out of range

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_param

int mx240_param (int *frequenceTalkerDatalog, int *numberofDatalogelements);

Purpose

Parameters: return DataLog informations from the

instrument.

Parameter List

frequenceTalkerDatalog

Variable Type       int (passed by reference)

Displays the timer value for Datalog or Talker

measurement:

Variable Type:  integer

Variable Range:  0  - 0.5 second

1  -   1 second

2  -   2 seconds

3  -   5 seconds

4  -  10 seconds

5  -  20 seconds

6  -  30 seconds

7  -  40 seconds

8  -  50 seconds

9  -   1 minute

10 -   2 minutes

11 -   3 minutes

12 -   4 minutes

13 -   5 minutes

numberofDatalogelements

Variable Type       int (passed by reference)

Displays the number of Datalog elements:

Variable Type:  integer

Variable Range:  0  to 512

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_power

int mx240_power (int unity, int _2ndFunction, int range, int compen,

int adjustementofCapacity, double referenceF);

Purpose

Switches the instrument to power measurement.

Parameter List

unity

Variable Type       int

Specifies the type of measurement on first display.

Variable Type: integer

Valid Range:   3 -  VA

4 -  W

5 -  var

Default Value: 3 -  VA

_2ndFunction

Variable Type       int

Specifies the type of measurement on second display.

Variable Type: integer

Valid Range:   0 -  None

2 -  Cos

3 -  Capacity

4 -  V AC+DC

8 -  A AC+DC

Default Value: 0 -  None

range

Variable Type       int

Specifies the range of the measurement

Variable Type:  integer

Variable Range:  8 - Auto Range

0 - 2.000k or 20.00k

1 - 20.00k or 200.0k

2 - 200.0k or 2000 k

Default Value:   8 - Auto Range

compen

Variable Type       int

Specifies the mode of Compensation of Capacity

Variable Type:  integer

Variable Range:  0 - Off

1 - On

Default Value:   0

adjustementofCapacity

Variable Type       int

Specifies if the capacity value is the value stored in the field

'Referenve (nF)'

Variable Type:  integer

Variable Range:  0 - Off

1 - On

Default Value:   0

referenceF

Variable Type       double

Specifies the adjust capacity value in F.

Variable Type:   double

Variable Range:    0.0 to  99999.0 mF

Default Value:   0.0

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

-1   First Function out of range

-2   Range out of range

-3   Second Function out of range

-4   Compensation Mode out of range

-5   Adjust Mode out of range

-6   Capacity value out of range

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

300   Function not available

301   Wrong position of the rotary switch

302   Two arguments are incompatible

303   Measurement is overloaded (OL)

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_resistance

int mx240_resistance (int range, int continuity);

Purpose

Switches the instrument to resistance measurement.

Parameter List

range

Variable Type       int

Specifies the range of the measurement

Variable Type:  integer

Variable Range:  8 - Auto Range

0 - 200.0 Ohm

1 - 2000  Ohm

Default Value:   8 - Auto Range

continuity

Variable Type       int

Specifies the mode of the measurement

Variable Type:  integer

Variable Range:  0 - Ohm

1 - Continuity

Default Value:   0 - Ohm

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

-2   Range out of range

-3   Continuity out of range

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

300   Function not available

301   Wrong position of the rotary switch

302   Two arguments are incompatible

303   Measurement is overloaded (OL)

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_status

int mx240_status (int *instrument, double *version, int *rotaryswitchhard,

int *rotaryswitchsoft, int *typeoffirstDisplay,

int *typeofsecondDisplay, int *lastButton);

Purpose

Configuration: return data informations from the

instrument.

Parameter List

instrument

Variable Type       int (passed by reference)

Displays the reference of the instrument:

0 or 4

Variable Type : integer

version

Variable Type       double (passed by reference)

Displays the reference of the instrument:

XX.XX

Variable Type : double

rotaryswitchhard

Variable Type       int (passed by reference)

Displays the position of the rotary switch.

1 => A

2 => V

3 => VA

4 => W

5 => var

6 => Wh

7 => OHM

Variable Type: integer

rotaryswitchsoft

Variable Type       int (passed by reference)

Displays the position of the rotary switch.

1 => A

2 => V

3 => VA

4 => W

5 => var

6 => Wh

7 => OHM

Variable Type: integer

typeoffirstDisplay

Variable Type       int (passed by reference)

Displays the configuration of the measurement.

10 => A AC+DC   20 => V AC+DC   30 => Cos

11 => A AC      21 => V AC      34 => VA

12 => A DC      22 => V DC      44 => Watt

13 => CF A      23 => CF V      54 => var

64 => Wh

74 => Ohm

82 => Continuity

Variable Type: integer

typeofsecondDisplay

Variable Type       int (passed by reference)

Displays the configuration of the measurement.

0  => none       10 => A AC+DC    20 => V AC+DC

1  => Hz         11 => A AC       21 => V AC

3  => Capacity   12 => A DC       22 => V DC

30 => Cos

Variable Type: integer

lastButton

Variable Type       int (passed by reference)

Displays the last button pressed before remote control

0 =>  First Display

1 =>  Minus

2 =>  Hold

3 =>  Range

4 =>  Menu

5 =>  Second Display

6 =>  Cal.

7 =>  Plus

8 =>  Surv.

Variable Type: integer

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_surveillance

int mx240_surveillance (int surv);

Purpose

Specifies the SURV mode.

Parameter List

surv

Variable Type       int

Specifies the SURV mode to access the surveillance

mode in which minimum, maximum ,sliding average and peak

values of the current measurement are stored.

Variable Type:  integer

Variable Range:  0 - Off

1 - On

Default Value:   0

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

-1   Timer value out of range

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_talker

int mx240_talker (int talker);

Purpose

Specifies the TALKER mode.

Parameter List

talker

Variable Type       int

Specifies the Talker mode to send the current value.

Variable Type:  integer

Variable Range:  0 - Off

1 - On

Default Value:   0

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

-1   Timer value out of range

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_timecos

int mx240_timecos (int timer, double cosinus);

Purpose

Changes the Timer value of the instrument. It is

used for talker and datalog measurement.

Changes the Cosinus value of the instrument. It is

used for Capacity Adjustement.

Parameter List

timer

Variable Type       int

Specifies the timer value for Datalog or Talker

measurement:

Variable Type:  integer

Variable Range:  0  - 0.5 second

1  -   1 second

2  -   2 seconds

3  -   5 seconds

4  -  10 seconds

5  -  20 seconds

6  -  30 seconds

7  -  40 seconds

8  -  50 seconds

9  -   1 minute

10 -   2 minutes

11 -   3 minutes

12 -   4 minutes

13 -   5 minutes

Default Value:   0  - 0.5 second

cosinus

Variable Type       double

Specifies the Cosinus value.

Variable Type:   double

Variable Range: 0.70 to 1.00

Default Value:   1.00

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

-1   Timer value out of range

-2   Cosinus value out of range

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

4xx   See RS232 Library error codes

(rs232err = 400 - Error)

________________________________________________________________________________

mx240_voltage

int mx240_voltage (int _1stFunction, int range, int _2ndFunction);

Purpose

Switches the instrument to voltage measurement.

Parameter List

_1stFunction

Variable Type       int

Specifies the type of measurement on first display.

Variable Type: integer

Valid Range:   0 -  AC+DC

1 -  AC

2 -  DC

3 -  CF

Default Value: 0 -  AC+DC

range

Variable Type       int

Specifies the range of the measurement

Variable Type:  integer

Variable Range:  8 - Auto Range

0 - 400 V

1 - 750 V

Default Value:   8 - Auto Range

_2ndFunction

Variable Type       int

Specifies the type of measurement on second display.

Variable Type: integer

Valid Range:   0 -  None

1 -  Hz

4 -  V AC+DC

8 -  A AC+DC

9 -  A AC

10-  A DC

Default Value: 0 -  None

Return Value

Displays the results of the instrument driver

function call:

Variable Type: Integer

Error Codes:

0   Success

-1   First Function out of range

-2   Range out of range

-3   Second Function out of range

220   Unable to open instrument

223   Instrument identification query failed

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

300   Function not available

301   Wrong position of the rotary switch

302   Two arguments are incompatible

303   Measurement is overloaded (OL)

304   Incompatible with SURV mode

4xx   See RS232 Library error codes

(rs232err = 400 - Error)