Metrix OX2000 DSO 200MS/s (VISA)

________________________________________________________________________________

Introduction:

This instrument driver provides programming support for Metrix OX2000 DSO 200MS/s (VISA).

It contains functions for opening, configuring, taking measurements from, and

closing the instrument.

________________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

For GPIB instrument drivers:

-  the instrument is connected to the GPIB.

-  the GPIB address supplied to the initialize function must match the

GPIB address of the instrument.

For RS-232 instrument drivers:

-  the instrument is connected to the RS-232 interface.

-  the COM port, baud rate, parity, and timeout supplied to the

initialize function must match the settings of the instrument.

________________________________________________________________________________

Error and Status Information:

Each function in this instrument driver returns a status code that

either indicates success or describes an error or warning condition.

Your program should examine the status code from each call to an

instrument driver function to determine if an error occurred.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

The description of each instrument driver function lists possible error

codes and their meanings

________________________________________________________________________________

How To Use This Document:

Use this document as a programming reference manual.

It describes each function in the Metrix OX2000 DSO 200MS/s (VISA)

instrument driver.  The functions appear in alphabetical order, with a description

of the function and its C syntax, a description of each parameter, and a list

of possible error codes.

________________________________________________________________________________

Function Tree Layout:

Metrix OX2000 DSO 200MS/s (VISA)       Function Name:

Initialize                                  mxOX2000_init

Application Functions

Example                                  mxOX2000_example

Configure Functions

Vertical Setup                           mxOX2000_verticalSetup

Horizontal

Horizontal Setup                      mxOX2000_horizontalSetup

Horizontal Display                    mxOX2000_horizontalDisplay

Trigger

General Trigger                       mxOX2000_setGeneralTrigger

Edge Trigger                          mxOX2000_setEdgeTrigger

Delayed Trigger                       mxOX2000_setDelayedTrigger

Count Trigger                         mxOX2000_setCountTrigger

Video Trigger                         mxOX2000_setVideoTrigger

Display

Options                               mxOX2000_setOptionDisplay

Waveform                              mxOX2000_setWaveformDisplay

Functions                             mxOX2000_setFunction

Measurements                          mxOX2000_setMeasurement

Cursors                               mxOX2000_setCursors

Hardcopy Setup                           mxOX2000_setHardcopy

Action/Status Functions

Autoset                                  mxOX2000_runAutoset

Run/Stop                                 mxOX2000_runAcquisition

Hardcopy                                 mxOX2000_runHardcopy

Data

Read Cursors                             mxOX2000_readCursors

Read Measurements                        mxOX2000_readMeasurement

Waveform

Read Channel To Array                 mxOX2000_readChToArr

Read Channel To File                  mxOX2000_readChToFile

Read File To Array                    mxOX2000_readFileToArr

Write Array To Memory                 mxOX2000_writeArrToMem

Write File To Memory                  mxOX2000_writeFileToMem

Write Array To File                   mxOX2000_writeArrToFile

Transfer

Save/Recall Setup To/From File        mxOX2000_saveSetup

File Transfer                         mxOX2000_fileTransfer

Utility Functions

Write To Instrument                      mxOX2000_writeInstrData

Read Instrument Data                     mxOX2000_readInstrData

Reset                                    mxOX2000_reset

Self-Test                                mxOX2000_selfTest

Error-Query                              mxOX2000_errorQuery

Error Message                            mxOX2000_errorMessage

Revision Query                           mxOX2000_revisionQuery

Front Panel Lockout                      mxOX2000_frontPanLock

Close                                       mxOX2000_close

________________________________________________________________________________

Metrix OX2000 DSO 200MS/s (VISA)

This instrument module provides programming support for the Metrix OX2000 DSO

200MS/s (VISA).  The module is divided into the following functions:

Functions/Classes:

(1) Initialize:

This function initializes the instrument and sets it to a default

configuration.

(2) Application Functions: (Class)

This class contains high-level, test and measurement routines.  These

examples call other instrument driver functions to configure, start, and

read from the instrument.

(3) Configuration Functions: (Class)

This class of functions configures the instrument by setting acquisition

and system configuration parameters.

(4) Action/Status Functions: (Class)

This class of functions begins or terminates an acquisition. It also

provides functions which allow the user to determine the current status

of the instrument.

(5) Data Functions: (Class)

This class of functions transfers data to or from the instrument.

(6) Utility Functions: (Class)

This class of functions provides lower level functions to communicate

with the instrument, and change instrument parameters.

(7) Close:

This function takes the instrument offline.

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

mxOX2000_close

ViStatus mxOX2000_close (ViSession instrumentHandle);

Purpose

This function performs the following operations:

viClose (instrSession) and viClose (rmSession).

Notes:

(1) The instrument must be reinitialized to use it again.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_errorMessage

ViStatus mxOX2000_errorMessage (ViSession instrumentHandle, ViStatus errorCode,

ViChar errorMessage[]);

Purpose

This function takes the Status Code returned by the instrument driver

functions, interprets it and returns it as a user readable string.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  VI_NULL

errorCode

Variable Type       ViStatus

This control accepts the Status Code returned from the instrument

driver functions.

Default Value:

0 - VI_SUCCESS

errorMessage

Variable Type       ViChar[]

This control returns the interpreted Status Code as a user readable

message string.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFF0085  Unknown status.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_errorQuery

ViStatus mxOX2000_errorQuery (ViSession instrumentHandle, ViInt32 *errorCode,

ViChar errorMessage[]);

Purpose

This function reads an error code from the instrument's error queue.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

errorCode

Variable Type       ViInt32 (passed by reference)

This control returns the error code read from the instrument's error

queue.

errorMessage

Variable Type       ViChar[]

This control returns the error message string read from the

instrument's error message queue.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_example

ViStatus mxOX2000_example (ViSession instrumentHandle, ViInt16 source,

ViInt16 timebaseA, ViInt16 voltageRange,

ViInt16 coupling, ViReal64 waveformArray[],

ViInt32 *numberofsamples, ViReal64 *XStart,

ViReal64 *XIncrement);

Purpose

This function is an example of this driver's functions use.

The instrument is set to the single shot mode.

User can change the source channel, the timebase, the input voltage range

and the input coupling.

A trigger is then sent to the instrument.

The function waits until the acquisition is completed.

Channel is read out from the instrument.

Waveform is interpolated if needed.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

source

Variable Type       ViInt16

This control selects the source channel.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

timebaseA

Variable Type       ViInt16

This control selects the main timebase (A) value.

Valid Values:

0  - 2ns

1  - 5ns

2  - 10ns

3  - 20ns

4  - 50ns

5  - 100ns

6  - 200ns

7  - 500ns - (Default Value)

8  - 1s

9  - 2s

10 - 5s

11 - 10s

12 - 20s

13 - 50s

14 - 100s

15 - 200s

16 - 500s

17 - 1ms

18 - 2ms

19 - 5ms

20 - 10ms

21 - 20ms

22 - 50ms

23 - 100ms

24 - 200ms

25 - 500ms

26 - 1s

27 - 2s

28 - 5s

29 - 10s

30 - 20s

31 - 50s

voltageRange

Variable Type       ViInt16

This control sets vertical sense for the specified channel.

Valid Values:

0  - 2mV/div

1  - 5mV/div

2  - 10mV/div

3  - 20mV/div

4  - 50mV/div

5  - 100mV/div

6  - 200mV/div

7  - 500mV/div

8  - 1V/div

9  - 2V/div

10 - 5V/div

11 - 10V/div  (Default Value)

12 - 20V/div

13 - 50V/div

14 - 100V/div

15 - 200V/div

16 - 500V/div

17 - 1kV/div

Note : This control value must be consistent with the probe value.

See the Digiscope 2000 Operating Instructions manual for more

details. If it is inconsistent an error is generated :

VI_ERROR_INCONSISTENT_PARAMETER7 (0xBFFC0907)

coupling

Variable Type       ViInt16

This control selects the input coupling of the selected channel.

Valid Values:

0 - AC  (Default Value)

1 - DC

2 - GND

waveformArray

Variable Type       ViReal64[]

This control specifies the name of the array in which to place the

waveform data.

Notes:

(1) The Real Array must be large enough to hold the requested points.

numberofsamples

Variable Type       ViInt32 (passed by reference)

This control returns the number of valid points read from the

selected channel.

XStart

Variable Type       ViReal64 (passed by reference)

This control returns the time of the first data point in seconds

(with respect to the trigger point).

XIncrement

Variable Type       ViReal64 (passed by reference)

This control returns the incrementing value in seconds for the X

axis.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Channel) out of range.

BFFC0003  Parameter 3 (Timebase) out of range.

BFFC0004  Parameter 4 (Voltage Range) out of range.

BFFC0005  Parameter 5 (Coupling) out of range.

BFFC0803  Error scanning response.

BFFC0900  Instrument specific error.

BFFC090C  Channel is not active.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_fileTransfer

ViStatus mxOX2000_fileTransfer (ViSession instrumentHandle, ViString fileName,

ViInt16 sourceDrive, ViBoolean overwriteFile,

ViInt16 targetDrive);

Purpose

This function performs a file transfer.

Note :

Size of files to be transfered from Computer to OX2000 must be less than

65498 bytes.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

fileName

Variable Type       ViString

This function specifies the name of the data file.

Default Value: ""

Notes:

(1) When reading a file from Computer, precede all backslashes with

an additional backslash, e.g. if the file is in the \cvi\instr

directory, enter the filename as "\\cvi\\instr\\anyfile".

(2) File name must be 8-character long + ".???" extension.

If not an error is generated :0xBFFC0002

(3) If the file already exists on the target drive and overwrite is

not selected an error is generated : 0xBFFC090E

(4) If file does not exist on the source drive an error is generated

: 0xBFFC090D

sourceDrive

Variable Type       ViInt16

This control selects the source drive.

Valid Values:

0 - Computer  (Default Value)

1 - OX2000 Memory

2 - OX2000 PC-CARD

overwriteFile

Variable Type       ViBoolean

This function specifies whether or not to write over the specified

file if it already exists on disk.

Valid Range:

VI_OFF (0) - No  (Default Value)

VI_ON  (1) - Yes

targetDrive

Variable Type       ViInt16

This control selects the target drive.

Valid Values:

0 - Computer

1 - OX2000 Memory  (Default value)

2 - OX2000 PC-CARD

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Filename) out of specifications.

BFFC0003  Parameter 3 (Source Drive) out of range.

BFFC0004  Parameter 4 (Overwrite File?) out of range.

BFFC0005  Parameter 5 (Target Drive) out of range.

BFFC0800  Error opening file.

BFFC0801  Error writing to file.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFC090D  File not found.

BFFC090E  File already exists.

BFFC0910  Invalid file content.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_frontPanLock

ViStatus mxOX2000_frontPanLock (ViSession instrumentHandle, ViBoolean mode);

Purpose

This function controls the instrument's front panel lockout state.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

mode

Variable Type       ViBoolean

Controls the front panel lockout.

Valid Range:

VI_OFF (0) - Unlock  (Default Value)

VI_ON  (1) - Lock

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Mode) out of range.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_horizontalDisplay

ViStatus mxOX2000_horizontalDisplay (ViSession instrumentHandle, ViInt16 trace,

ViInt16 x, ViInt16 y, ViInt16 roll,

ViBoolean cumul);

Purpose

This function configures the horizontal display of the instrument.

The following settings are available :

- Trace

- X (for XY mode)

- Y (for XY mode)

- Roll mode

- Cumul mode

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

trace

Variable Type       ViInt16

This control selects the traces to be displayed.

Valid Values:

0 - A  (Default Value)

1 - A

2 - A&B

3 - XY

4 - A&XY

x

Variable Type       ViInt16

This control selects the X input for XY mode.

Valid Values:

1 - 1  (Default Value)

2 - 2

3 - 3

4 - 4

y

Variable Type       ViInt16

This control selects the Y input for XY mode.

Valid Values:

1 - 1  (Default Value)

2 - 2

3 - 3

4 - 4

roll

Variable Type       ViInt16

This control selects the roll mode.

Valid Values:

0 - Off (Default Value)

1 - On

2 - Free

cumul

Variable Type       ViBoolean

This control selects the cumul mode.

Valid Values:

VI_OFF (0) - Off (Default Value)

VI_ON  (1) - On

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Trace) out of range.

BFFC0003  Parameter 3 (X) out of range.

BFFC0004  Parameter 4 (Y) out of range.

BFFC0005  Parameter 5 (Roll) out of range.

BFFC0006  Parameter 6 (Cumul) out of range.

BFFC0803  Error scanning response.

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_horizontalSetup

ViStatus mxOX2000_horizontalSetup (ViSession instrumentHandle, ViInt16 size,

ViInt16 pretrigger, ViReal64 runafterdelays,

ViInt16 rangeA, ViReal64 delayAs,

ViInt16 rangeB, ViReal64 delayBs);

Purpose

This function configures the horizontal settings of the instrument.

The following settings are available :

- Acquisition size

- Pretrigger

- Run after delay

- RangeA (the main timebase)

- DelayA (the offset on the main timebase)

- RangeB (the auxiliary timebase)

- DelayB (the offset on the auxiliary timebase)

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

size

Variable Type       ViInt16

This control selects the acquisition size.

Valid Values:

0 - 1000  (Default Value)

1 - 2000

2 - 5000

3 - 10000

pretrigger

Variable Type       ViInt16

This control selects the pretrigger position.

Valid Values:

0  - 0%

1  - 10%

2  - 20%

3  - 30%

4  - 40%

5  - 50% - (Default Value)

6  - 60%

7  - 70%

8  - 80%

9  - 90%

10 - 100%

runafterdelays

Variable Type       ViReal64

This control selects the Run After Delay value (in seconds).

Default Value: 0.0

rangeA

Variable Type       ViInt16

This control selects the main timebase (A) value.

Valid Values:

0  - 2ns

1  - 5ns

2  - 10ns

3  - 20ns

4  - 50ns

5  - 100ns

6  - 200ns

7  - 500ns - (Default Value)

8  - 1s

9  - 2s

10 - 5s

11 - 10s

12 - 20s

13 - 50s

14 - 100s

15 - 200s

16 - 500s

17 - 1ms

18 - 2ms

19 - 5ms

20 - 10ms

21 - 20ms

22 - 50ms

23 - 100ms

24 - 200ms

25 - 500ms

26 - 1s

27 - 2s

28 - 5s

29 - 10s

30 - 20s

31 - 50s

delayAs

Variable Type       ViReal64

This control selects the Offset delay value (in seconds)

for the main timebase (A).

Default Value: 0.0

rangeB

Variable Type       ViInt16

This control selects the auxiliary timebase (B) value.

Valid Values:

0  - 2ns

1  - 5ns

2  - 10ns

3  - 20ns

4  - 50ns

5  - 100ns

6  - 200ns

7  - 500ns - (Default Value)

8  - 1s

9  - 2s

10 - 5s

11 - 10s

12 - 20s

13 - 50s

14 - 100s

15 - 200s

16 - 500s

17 - 1ms

18 - 2ms

19 - 5ms

20 - 10ms

21 - 20ms

22 - 50ms

23 - 100ms

24 - 200ms

25 - 500ms

26 - 1s

27 - 2s

28 - 5s

29 - 10s

30 - 20s

31 - 50s

32 - 100s

33 - 200s

34 - 500s

Note : This control value must be consistent with the main timebase

(A) and the acquisition size values. See the Digiscope 2000 Operating

Instructions manual for more details. If it is inconsistent an error

is generated : VI_ERROR_INCONSISTENT_PARAMETER7 (0xBFFC0907)

delayBs

Variable Type       ViReal64

This control selects the Offset delay value (in seconds)

for the auxiliary timebase (B).

Default Value: 0.0

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Size) out of range.

BFFC0003  Parameter 3 (Pretrig) out of range.

BFFC0004  Parameter 4 (Delay) out of range.

BFFC0005  Parameter 5 (RangeA) out of range.

BFFC0006  Parameter 6 (DelayA) out of range.

BFFC0007  Parameter 7 (RangeB) out of range.

BFFC0008  Parameter 8 (DelayB) out of range.

BFFC0803  Error scanning response.

BFFC0900  Instrument specific error.

BFFC0907  Parameter 7 (RangeB) is inconsistent.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_init

ViStatus mxOX2000_init (ViRsrc resourceName, ViBoolean IDQuery,

ViBoolean resetDevice, ViSession *instrumentHandle);

Purpose

This function performs the following initialization actions:

- Opens a session to the Default Resource Manager resource and a session

to the specified device using the interface and address specified in the

Resource_Name control.

- Performs an identification query on the Instrument.

- Resets the instrument to a known state.

- Serial Parameters are: Speed = 19200, Parity = None, Data = 8 bits,

Protocol = RTS/CTS (hard)

- Sends initialization commands to the instrument that set any necessary

programmatic variables such as Headers Off, Short Command form, and Data

Transfer Binary to the state necessary for the operation of the

instrument driver.

- Returns an Instrument Handle which is used to differentiate between

different sessions of this instrument driver.

- Each time this function is invoked a Unique Session is opened.  It is

possible to have more than one session open for the same resource.

Parameter List

resourceName

Variable Type       ViRsrc

This control specifies the interface and address of the device that

is to be initialized (Instrument Descriptor). The exact grammar to be

used in this control is shown in the note below.

Default Value:  "GPIB::1"

Notes:

(1) Based on the Instrument Descriptor, this operation establishes a

communication session with a device.  The grammar for the Instrument

Descriptor is shown below.  Optional parameters are shown in square

brackets ([]).

Interface   Grammar

------------------------------------------------------

GPIB        GPIB[board]::primary address[::secondary address]

[::INSTR]

Serial      ASRL<port>::INSTR

The GPIB keyword is used with GPIB instruments.

The ASRL keyword is used with Serial instruments.

The default value for optional parameters are shown below.

Optional Parameter          Default Value

-----------------------------------------

board                       0

secondary address           none - 31

port                        1 - 16

IDQuery

Variable Type       ViBoolean

This control specifies if an ID Query is sent to the instrument

during the initialization procedure.

Valid Range:

VI_OFF (0) - Skip Query

VI_ON  (1) - Do Query (Default Value)

Notes:

(1) Under normal circumstances the ID Query ensures that the

instrument initialized is the type supported by this driver. However

circumstances may arise where it is undesirable to send an ID Query

to the instrument.  In those cases; set this control to "Skip Query"

and this function will initialize the selected interface, without

doing an ID Query.

resetDevice

Variable Type       ViBoolean

This control specifies if the instrument is to be reset to its

power-on settings during the initialization procedure.

Valid Range:

VI_OFF (0) - Don't Reset

VI_ON  (1) - Reset Device (Default Value)

Notes:

(1) If you do not want the instrument reset. Set this control to

"Don't Reset" while initializing the instrument.

instrumentHandle

Variable Type       ViSession (passed by reference)

This control returns an Instrument Handle that is used in all

subsequent function calls to differentiate between different sessions

of this instrument driver.

Notes:

(1) Each time this function is invoked a Unique Session is opened.

It is possible to have more than one session open for the same

resource.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (ID Query) out of range.

BFFC0003  Parameter 3 (Reset Device) out of range.

BFFC0011  Instrument returned invalid response to ID Query

BFFC0803  Error Scanning Response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_readChToArr

ViStatus mxOX2000_readChToArr (ViSession instrumentHandle, ViInt16 source,

ViBoolean interpolation,

ViReal64 waveformArray[],

ViInt32 *numberofsamples, ViReal64 *XStart,

ViReal64 *XIncrement);

Purpose

This function reads the specified channel from the instrument and places

it in an array.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

source

Variable Type       ViInt16

This control selects the source channel.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

interpolation

Variable Type       ViBoolean

This control selects whether to interpolate or not the samples when

in ETS acquisition (see OX2000 operating instructions manual for more

details).

Valid Values:

VI_OFF (0) - Off (Default Value)

VI_ON  (1) - Linear

Note

(1) If interpolation is not selected, invalid points take the

following value : 1E+100

waveformArray

Variable Type       ViReal64[]

This function specifies the name of array in which to place the

waveform data.

Notes:

(1) The Real Array must be large enough to hold the requested points.

numberofsamples

Variable Type       ViInt32 (passed by reference)

This function returns the number of valid points read from the

selected channel.

XStart

Variable Type       ViReal64 (passed by reference)

This function returns the time of the first data point in seconds

(with respect to the trigger point).

XIncrement

Variable Type       ViReal64 (passed by reference)

This function returns the incrementing value in seconds for the X

axis.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Source) out of range.

BFFC0003  Parameter 3 (Interpolation) out of range.

BFFC0803  Error scanning response.

BFFC0900  Instrument specific error.

BFFC090C  Channel is not active.

BFFC0910  Invalid file content.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_readChToFile

ViStatus mxOX2000_readChToFile (ViSession instrumentHandle, ViInt16 source,

ViInt16 drives, ViString fileName,

ViBoolean overwriteFile);

Purpose

This function reads a channel from the instrument and stores it as a

file.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

source

Variable Type       ViInt16

This control selects the source channel.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

drives

Variable Type       ViInt16

This control selects the destination drive.

Valid Values:

0 - Computer

1 - OX2000 Memory

2 - OX2000 PC-CARD

fileName

Variable Type       ViString

This function specifies the name of the wave data file.

Default Value: ""

Notes:

(1) When storing a file to Computer, precede all backslashes with an

additional backslash, e.g. if the file is in the \cvi\instr

directory, enter the filename as "\\cvi\\instr\\anyfile".

(2) File name must be 8-character long plus ".TRC" extension.

If not, an error is generated :0xBFFC0004

(3) If the target file already exists and overwrite is not selected,

an error is generated : 0xBFFC090E

overwriteFile

Variable Type       ViBoolean

This function specifies whether or not to write over the specified

file if it already exists.

Valid Range:

VI_OFF (0) - No  (Default Value)

VI_ON  (1) - Yes

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Source) out of range.

BFFC0003  Parameter 3 (Drive) out of range.

BFFC0004  Parameter 4 (Filename) out of specifications.

BFFC0005  Parameter 5 (Overwrite File?) out of range.

BFFC0801  Error writing to file.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFC090C  Channel is not active.

BFFC090E  File already exists.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_readCursors

ViStatus mxOX2000_readCursors (ViSession instrumentHandle, ViInt16 reference,

ViReal64 *t1s, ViReal64 *t2s, ViReal64 *dTs,

ViReal64 *v1v, ViReal64 *v2v, ViReal64 *dVv);

Purpose

This function reads cursors positions.

All measured quantities are in their basic units.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

reference

Variable Type       ViInt16

This control selects the reference channel for the cursors positions.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

t1s

Variable Type       ViReal64 (passed by reference)

This control outputs the position of the vertical cursor #1.

t2s

Variable Type       ViReal64 (passed by reference)

This control outputs the position of the vertical cursor #2.

dTs

Variable Type       ViReal64 (passed by reference)

This control outpus the difference between vertical cursors.

v1v

Variable Type       ViReal64 (passed by reference)

This control outpus the position of the horizontal cursor #1.

v2v

Variable Type       ViReal64 (passed by reference)

This control outpus the position of the horizontal cursor #2.

dVv

Variable Type       ViReal64 (passed by reference)

This control outpus the difference between horizontal cursors.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Reference) out of range.

BFFC0003  Parameter 3 (T1) out of range.

BFFC0004  Parameter 4 (T2) out of range.

BFFC0005  Parameter 5 (dT) out of range.

BFFC0006  Parameter 6 (V1) out of range.

BFFC0007  Parameter 7 (V2) out of range.

BFFC0008  Parameter 8 (dV) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_readFileToArr

ViStatus mxOX2000_readFileToArr (ViSession instrumentHandle, ViInt16 drives,

ViString fileName, ViBoolean interpolation,

ViReal64 waveformArray[],

ViInt32 *numberofsamples, ViReal64 *XStart,

ViReal64 *XIncrement);

Purpose

This function reads a stored waveform file and places the data in an

array.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

drives

Variable Type       ViInt16

This control selects the source drive.

Valid Values:

0 - Computer

1 - OX2000 Memory

2 - OX2000 PC-CARD

fileName

Variable Type       ViString

This function specifies the name of the wave data file.

Default Value: ""

Notes:

(1) When reading a file from Computer, precede all backslashes with

an additional backslash, e.g. if the file is in the \cvi\instr

directory, enter the filename as "\\cvi\\instr\\anyfile".

(2) File name must be 8-character long + ".TRC" extension.

If not an error is generated :0xBFFC0003

(3) If file does not exist an error is generated : 0xBFFC090D

interpolation

Variable Type       ViBoolean

This control selects whether to interpolate or not the samples when

in ETS acquisition (see OX2000 operating instructions manual for more

details).

Valid Values:

VI_OFF (0) - Off (Default Value)

VI_ON  (1) - Linear

Note

(1) If interpolation is not selected, invalid points take the

following value : 1E+100

waveformArray

Variable Type       ViReal64[]

This function specifies the name of array in which to place the

waveform data.

Notes:

(1) The Real Array must be large enough to hold the requested points.

numberofsamples

Variable Type       ViInt32 (passed by reference)

This function returns the number of valid points read from the

selected file.

XStart

Variable Type       ViReal64 (passed by reference)

This function returns the time of the first data point in seconds

(with respect to the trigger point).

XIncrement

Variable Type       ViReal64 (passed by reference)

This function returns the incrementing value in seconds for the X

axis.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Drive) out of range.

BFFC0003  Parameter 3 (Filename) out of specifications.

BFFC0004  Parameter 4 (Interpolation) out of range.

BFFC0800  Error opening file.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFC090D  File not found.

BFFC0910  Invalid file content.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_readInstrData

ViStatus mxOX2000_readInstrData (ViSession instrumentHandle,

ViInt32 numberBytesToRead, ViChar readBuffer[],

ViInt32 *numBytesRead);

Purpose

This function reads data from the instrument's output buffer and returns

it to the specified variable in memory.

Notes:

(1) Because the instrument may return both numeric and text data in

response to queries this function returns the data in string format.

(2) If valid data is not available at the instrument's output buffer when

this function is called the instrument will hang up and the function will

not return until it times out.  If the time-out is disabled this function

will hang indefinitely and it may be necessary to reboot the computer to

break out.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

numberBytesToRead

Variable Type       ViInt32

The number of bytes specified by this control should be greater than

or equal to the number of bytes which are to be read from the

instrument. If the actual number of bytes to be read is greater than

the number this control specifies then multiple reads will be

required to empty the instrument's output buffer.

If the instrument's output buffer is not emptied the instrument may

return invalid data when future reads are performed.

Default Value:  50 (See NOTE 2)

Notes:

(1) If NO DATA is available at the instrument's output buffer when

this function is called the instrument will hang up and the function

will not return until it times out.  If the time-out is disabled,

this function will hang indefinitely and it may be necessary to

reboot the computer.

(2) If the number of bytes expected is greater than 50 the value of

this control may be increased. If the actual number of bytes read is

larger than the declared size of the read buffer a run-time error

will be generated.

readBuffer

Variable Type       ViChar[]

The incoming data from the instrument is placed into this variable.

Notes:

(1) Because the instrument may return both numeric and text data in

response to queries this function returns the data in string format.

(2) This function does not overwrite any old data left in the string

variable from the last time the function was called. String data in

LabWindows/CVI is terminated with an ASCII null(0x0) and string

manipulation functions will only recognize data before the ASCII

null.

(3) The declared size of the string variable must be greater than the

actual number of bytes read from the instrument if it is not a

run-time error will be generated.

numBytesRead

Variable Type       ViInt32 (passed by reference)

This variable contains the actual number of bytes read from the

instrument. This is the value which is returned by the read function.

Notes:

(1) If the actual number of bytes read is less than the number of

bytes specified in the Number Bytes To Read control then the output

buffer has probably been emptied. If the read function fails and the

number of bytes read is 0, the most probable cause for the failure is

there was no data available at the instrument's output buffer.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_readMeasurement

ViStatus mxOX2000_readMeasurement (ViSession instrumentHandle,

ViInt16 measurement, ViInt16 reference,

ViReal64 *reading);

Purpose

This function forces the instrument to perform selected measurement and

returns the result.

All measured quantities are in their basic units.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measurement

Variable Type       ViInt16

This control specifies the measurement function.

Valid Values:

0  - RMS Voltage     [V]  (Default Value)

1  - Amplitude       [V]

2  - Peak to Peak    [V]

3  - Average Voltage [V]

4  - High Voltage    [V]

5  - Low Voltage     [V]

6  - Maximum Voltage [V]

7  - Minimum Voltage [V]

8  - Overshoot       [%]

9  - Frequency       [Hz]

10 - Period          [s]

11 - Fall Time       [s]

12 - Rise Time       [s]

13 - Pos Pulse Width [s]

14 - Neg Pulse Width [s]

15 - Dutycycle       [%]

reference

Variable Type       ViInt16

This control selects the reference channel for the specified

measurement.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

reading

Variable Type       ViReal64 (passed by reference)

This control outputs the result of the measurement.

All measured quantities are output in their basic units.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Measurement) out of range.

BFFC0003  Parameter 3 (Reference) out of range.

BFFC0004  Parameter 4 (Reading) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_reset

ViStatus mxOX2000_reset (ViSession instrumentHandle);

Purpose

This function resets the instrument to a known state.

See the Digiscope 2000 Operating Instructions manual for more details.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_revisionQuery

ViStatus mxOX2000_revisionQuery (ViSession instrumentHandle,

ViChar instrumentDriverRevision[],

ViChar firmwareRevision[]);

Purpose

This function returns the revision numbers of the instrument driver and

instrument firmware, and tells the user what firmware the driver is

compatible with, for example:

This instrument driver's Revision Number is "Rev 1.1, 04/98, CVI 4.0.1"

and is compatible with Instrument Firmware Revision "1.6".

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

instrumentDriverRevision

Variable Type       ViChar[]

This control returns the Instrument Driver Software Revision.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

firmwareRevision

Variable Type       ViChar[]

This control returns the Instrument Firmware Revision.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_runAcquisition

ViStatus mxOX2000_runAcquisition (ViSession instrumentHandle, ViInt16 mode,

ViBoolean wait);

Purpose

This function controls the acquisition system of the instrument.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

mode

Variable Type       ViInt16

This control sets the acquisition system of the instrument.

Valid Values:

0 - Continuous  (Default Value)

1 - Single Shot

2 - Abort

wait

Variable Type       ViBoolean

This control selects whether to wait until a single shot operation is

completed.

Valid Values:

VI_OFF (0) - No  (Default Value)

VI_ON  (1) - Yes

Note:

This control is valid only if the single shot trigger mode is

selected.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Mode) out of range.

BFFC0003  Parameter 3 (Wait?) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_runAutoset

ViStatus mxOX2000_runAutoset (ViSession instrumentHandle);

Purpose

This function performs an autoset of the instrument.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0803  Error scanning response.

BFFC0900  Instrument specific error.

BFFC090F  No channel is enabled.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_runHardcopy

ViStatus mxOX2000_runHardcopy (ViSession instrumentHandle, ViBoolean action);

Purpose

This function controls the hardcopy system of the instrument.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

action

Variable Type       ViBoolean

This control selects whether to run the hardcopy or to abort it.

Valid Values:

VI_OFF (0) - Abort

VI_ON  (1) - Run     (Default Value)

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Action) out of range.

BFFC0803  Error scanning response

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_saveSetup

ViStatus mxOX2000_saveSetup (ViSession instrumentHandle, ViBoolean mode,

ViInt16 drives, ViString fileName,

ViBoolean overwriteFile);

Purpose

This function save or recall set-up to or from file.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

mode

Variable Type       ViBoolean

This control selects between Save and Recall mode.

Valid Range:

VI_OFF (0) - Recall

VI_ON  (1) - Save  (Default Value)

drives

Variable Type       ViInt16

This control selects the target/source drive.

Valid Values:

0 - Computer  (Default Value)

1 - OX2000 Memory

2 - OX2000 PC-CARD

fileName

Variable Type       ViString

This function specifies the name of the setup data file.

Default Value: ""

Notes:

(1) When reading/writing a file from/to Computer, precede all

backslashes with an additional backslash, e.g. if the file is in the

\cvi\instr directory, enter the filename as "\\cvi\\instr\\anyfile".

(2) File name must be 8-character long + ".CFG" extension.

If not an error is generated :0xBFFC0004

(3) When saving a setup : if the target file already exists and

overwrite is not selected an error is generated : 0xBFFC090E

(4) When recalling a setup : if file does not exist an error is

generated : 0xBFFC090D

overwriteFile

Variable Type       ViBoolean

This function specifies whether or not to write over the specified

file if it already exists on disk.

Valid Range:

VI_OFF (0) - No  (Default Value)

VI_ON  (1) - Yes

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Mode) out of range.

BFFC0003  Parameter 3 (Drive) out of range.

BFFC0004  Parameter 4 (Filename) out of specifications.

BFFC0005  Parameter 5 (Overwrite File?) out of range.

BFFC0800  Error opening file.

BFFC0801  Error writing to file.

BFFC0803  Error Scanning Response

BFFC0900  Instrument specific error.

BFFC090D  File not found.

BFFC090E  File already exists.

BFFC0910  Invalid file content.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_selfTest

ViStatus mxOX2000_selfTest (ViSession instrumentHandle, ViInt16 *selfTestResult,

ViChar selfTestMessage[]);

Purpose

This function runs the instrument's self test routine and returns the

test result(s).

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

selfTestResult

Variable Type       ViInt16 (passed by reference)

This control contains the value returned from the instrument self

test.  Zero means success.  For any other code, see the device's

operator's manual.

selfTestMessage

Variable Type       ViChar[]

This control contains the string returned from the self test. See the

device's operation manual for an explanation of the string's

contents.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0803  Error scanning response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_setCountTrigger

ViStatus mxOX2000_setCountTrigger (ViSession instrumentHandle,

ViBoolean triggerSource, ViInt16 qualifier,

ViInt16 counter, ViInt32 events);

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

triggerSource

Variable Type       ViBoolean

This control selects the trigger source.

Valid Values:

VI_OFF (0) - Qualifier  (Default Value)

VI_ON  (1) - Counter

qualifier

Variable Type       ViInt16

This control selects the qualifier source.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

5 - Line

counter

Variable Type       ViInt16

This control selects the counter source.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

5 - Line

events

Variable Type       ViInt32

This control sets the number of events to count.

Valid Range:

1 to 16 777 215

Default Value: 1

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Trigger source) out of range.

BFFC0003  Parameter 3 (Qualifier) out of range.

BFFC0004  Parameter 4 (Counter) out of range.

BFFC0005  Parameter 5 (Events) out of range.

BFFC0905  Parameter 5 (Events) is inconsistent.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_setCursors

ViStatus mxOX2000_setCursors (ViSession instrumentHandle, ViBoolean cursors,

ViInt16 reference, ViReal64 t1s, ViReal64 t2s,

ViReal64 v1V, ViReal64 v2V);

Purpose

This function sets vertical and horizontal cursors.

The following settings are available :

- Clear Cursors

- Reference

- T1

- T2

- V1

- V2

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

cursors

Variable Type       ViBoolean

This control allows you to clear or display the cursors.

Valid Values:

VI_OFF (0) - Off

VI_ON  (1) - On  (Default Value)

reference

Variable Type       ViInt16

This control selects the reference.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

t1s

Variable Type       ViReal64

This control sets the position of the vertical cursor #1.

Valid Range:

+/- 100 * horizSense (horizSense in s/div)

Default Value:

-2.5E-6 s

t2s

Variable Type       ViReal64

This control sets the position of the vertical cursor #2.

Valid Range:

+/- 100 * horizSense (horizSense in s/div)

Default Value:

-2.5E-6 s

v1V

Variable Type       ViReal64

This control sets the position of the horizontal cursor #1.

Valid Range:

+/- 80 * vertSense (vertSense in V/div)

Default Value:

40.0 V

v2V

Variable Type       ViReal64

This control sets the position of the horizontal cursor #2.

Valid Range:

+/- 80 * vertSense (vertSense in V/div)

Default Value:

40.0 V

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Clear Cursors) out of range.

BFFC0003  Parameter 3 (Reference) out of range.

BFFC0004  Parameter 4 (T1) out of range.

BFFC0005  Parameter 5 (T2) out of range.

BFFC0006  Parameter 6 (V1) out of range.

BFFC0007  Parameter 7 (V2) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_setDelayedTrigger

ViStatus mxOX2000_setDelayedTrigger (ViSession instrumentHandle,

ViInt16 triggerSource, ViInt16 qualifier,

ViReal64 delay);

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

triggerSource

Variable Type       ViInt16

This control selects the trigger source.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

5 - Line

qualifier

Variable Type       ViInt16

This control selects the qualifier source.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

5 - Line

delay

Variable Type       ViReal64

This control sets the delay value in seconds.

Valid Range:

40 ns  to  16 s

Default Value: 40 ns

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Trigger source) out of range.

BFFC0003  Parameter 3 (Qualifier) out of range.

BFFC0004  Parameter 4 (Delay) out of range.

BFFC0803  Error scanning response.

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_setEdgeTrigger

ViStatus mxOX2000_setEdgeTrigger (ViSession instrumentHandle,

ViInt16 triggerSource, ViBoolean window,

ViReal64 mainLevelvolts,

ViReal64 auxiliaryLevelvolts);

Purpose

This function configures the edge trigger mode of the instrument.

The following settings are available :

- Trigger source

- Window

- Main level

- Auxiliary level

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

triggerSource

Variable Type       ViInt16

This control selects the trigger source.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

5 - Line

window

Variable Type       ViBoolean

This control sets the window mode.

Valid Values:

VI_OFF (0) - Off

VI_ON  (1) - On  (Default Value)

Note : This control value must be consistent with the source value.

Window can't be set to 'On' with trigger source set to 'Line'. See

the Digiscope 2000 Operating Instructions manual for more details. If

it is inconsistent an error is generated :

VI_ERROR_INCONSISTENT_PARAMETER3 (0xBFFC0903)

mainLevelvolts

Variable Type       ViReal64

This control sets the main level of trigger for the trigger source.

Valid Range:

0.0 V  to  12 * vertSense (vertSense in Volts/div)

Default Value: 0.0 V

auxiliaryLevelvolts

Variable Type       ViReal64

This control sets the auxiliary level of trigger for the trigger

source in the window mode.

Valid Range:

0.0 V  to  12 * vertSense (vertSense in Volts/div)

Default Value: 0.0 V

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Source) out of range.

BFFC0003  Parameter 3 (Window) out of range.

BFFC0004  Parameter 4 (Main Level) out of range.

BFFC0005  Parameter 5 (Auxiliary Level) out of range.

BFFC0903  Parameter 3 (Window) is inconsistent.

BFFC0803  Error scanning response.

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_setFunction

ViStatus mxOX2000_setFunction (ViSession instrumentHandle, ViBoolean select,

ViInt16 function, ViInt16 source1,

ViInt16 source2);

Purpose

This function configures the mathematical facility of the instrument for

waveform operations.

The following settings are available :

- Select

- Function

- Source1

- Source2

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

select

Variable Type       ViBoolean

This control selects the function to configure.

Valid Values:

VI_OFF (0) - F1  (Default Value)

VI_ON  (1) - F2

function

Variable Type       ViInt16

This control selects the operation to perform.

Valid Values:

0 - Off  (Default Value)

1 - Add

2 - Sub

3 - Mul

source1

Variable Type       ViInt16

This control selects the first source.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

source2

Variable Type       ViInt16

This control selects the second source.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Select) out of range.

BFFC0003  Parameter 3 (Function) out of range.

BFFC0004  Parameter 4 (Source1) out of range.

BFFC0005  Parameter 5 (Source2) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_setGeneralTrigger

ViStatus mxOX2000_setGeneralTrigger (ViSession instrumentHandle,

ViInt16 triggerSource, ViBoolean slope,

ViInt16 coupling, ViBoolean HFRejection,

ViBoolean noiseRejection, ViInt16 type,

ViReal64 levelvolts,

ViReal64 holdoffseconds);

Purpose

This function configures the general trigger parameters of the

instrument.

The following settings are available :

- Trigger Source

- Slope

- Coupling

- HF Rejection

- Noise Rejection

- Type

- Level

- Holdoff

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

triggerSource

Variable Type       ViInt16

This control selects the trigger source to configure.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

slope

Variable Type       ViBoolean

This control sets the slope of the specified channel.

Valid Values:

VI_ON  (1) - Up (Default Value)

VI_OFF (0) - Down

coupling

Variable Type       ViInt16

This control selects the coupling mode of the specified channel.

Valid Values:

0 - AC (Default Value)

1 - DC

2 - LF Rej

HFRejection

Variable Type       ViBoolean

This control sets the HF Rejection of the specified channel.

Valid Values:

VI_OFF (0) - Off (Default Value)

VI_ON  (1) - On

Note : This control value must be consistent with the coupling value.

HF Rejection can't be set to 'On' with coupling set to 'LF

Rejection'. See the Digiscope 2000 Operating Instructions manual for

more details. If it is inconsistent an error is generated :

VI_ERROR_INCONSISTENT_PARAMETER5 (0xBFFC0905)

noiseRejection

Variable Type       ViBoolean

This control sets the noise rejection of the specified channel.

Valid Values:

VI_OFF (0) - Off  (Default Value)

VI_ON  (1) - On

type

Variable Type       ViInt16

This control selects the type of trigger.

Valid Values:

0 - Auto (Default Value)

1 - Triggered

2 - Peak-to-peak

levelvolts

Variable Type       ViReal64

This control sets level of trigger for the specified channel.

Valid Range:

0.0 V  to  12 * vertSense (vertSense in Volts/div)

Default Value: 0.0 V

holdoffseconds

Variable Type       ViReal64

This control sets the holdoff value.

Valid Range:

80.0 ns  to  1.0 s

Default Value: 80.0 ns

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Channel) out of range.

BFFC0003  Parameter 3 (Slope) out of range.

BFFC0004  Parameter 4 (Coupling) out of range.

BFFC0005  Parameter 5 (HF Rejection) out of range.

BFFC0006  Parameter 6 (Noise Rejection) out of range.

BFFC0007  Parameter 7 (Type) out of range.

BFFC0008  Parameter 8 (Level) out of range.

BFFC0803  Error scanning response.

BFFC0809  Parameter 9 (Holdoff) out of range.

BFFC0900  Instrument specific error.

BFFC0905  Parameter 5 (HF Rejection) is inconsistent.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_setHardcopy

ViStatus mxOX2000_setHardcopy (ViSession instrumentHandle, ViInt16 device,

ViInt16 port, ViString filename, ViBoolean type,

ViBoolean ofpages, ViInt16 pagenumber,

ViBoolean ofpens);

Purpose

This function sets hardcopy parameters.

The following settings are available :

- Device

- Port

- Filename

- # of Pages

- Page Number

- # of Pens

Note :

The '# of Pages' and 'Page Number' controls are only effective after

running a hardcopy with the 'runHardcopy' function.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

device

Variable Type       ViInt16

This control selects the device language.

Valid Values:

0  - IBM Proprinter (Default Value)

1  - EPSON ESC/P2

2  - HP-PCL

3  - HPGL

4  - Seiko I

port

Variable Type       ViInt16

This control selects the measurement to display.

Valid Values:

0  - File on CARD

1  - RS232 (Default Value)

2  - Centronics

3  - GPIB

filename

Variable Type       ViString

This control sets the filename where to store the hardcopy.

Default Value: ""

Note : It must be an 8-character name with no extension. Extension is

automatically added according to the device language. See the

Digiscope 2000 Operating Instructions manual for more details.

type

Variable Type       ViBoolean

This control sets the number of pages.

Valid Values:

VI_OFF (0) - Black & White  (Default Value)

VI_ON  (1) - Color

ofpages

Variable Type       ViBoolean

This control sets the number of pages.

Valid Values:

VI_OFF (0) - 1  (Default Value)

VI_ON  (1) - 4

Note : This control value must be consistent with the device value.

4-page printing mode is only valid with HPGL devices.

See the Digiscope 2000 Operating Instructions manual for more

details. If it is inconsistent an error is generated :

VI_ERROR_INCONSISTENT_PARAMETER6 (0xBFFC0906)

pagenumber

Variable Type       ViInt16

This control selects on which page to print the hardcopy.

Valid Values:

1 - 1  (Default Value)

2 - 2

3 - 3

4 - 4

Note : This parameter is only taken into account if 4-page printing

mode is active for an HPGL device.

ofpens

Variable Type       ViBoolean

This control sets the number of pens used.

Valid Values:

VI_OFF (0) - 2

VI_ON  (1) - 7 (Default Value)

Note : This parameter is only taken into account for an HPGL device.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Device) out of range.

BFFC0003  Parameter 3 (Port) out of range.

BFFC0004  Parameter 4 (Filename) out of range.

BFFC0005  Parameter 5 (Type) out of range.

BFFC0006  Parameter 6 (# of pages) out of range.

BFFC0007  Parameter 7 (Page Number) out of range.

BFFC0008  Parameter 8 (# of pens) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFC0906  Parameter 6 (# of pages) is inconsistent.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_setMeasurement

ViStatus mxOX2000_setMeasurement (ViSession instrumentHandle, ViInt16 edit,

ViInt16 measurement, ViInt16 reference);

Purpose

This function selects an automatical measurement from the instrument and

displays it.

The following settings are available :

- Edit

- Measurement

- Reference

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

edit

Variable Type       ViInt16

This control selects the edit mode.

Valid Values:

0 - Insert  (Default Value)

1 - Delete

2 - Clear all

measurement

Variable Type       ViInt16

This control selects the measurement to display.

Valid Values:

0  - RMS Voltage     [V]  (Default Value)

1  - Amplitude       [V]

2  - Peak to Peak    [V]

3  - Average Voltage [V]

4  - High Voltage    [V]

5  - Low Voltage     [V]

6  - Maximum Voltage [V]

7  - Minimum Voltage [V]

8  - Overshoot       [%]

9  - Frequency       [Hz]

10 - Period          [s]

11 - Fall Time       [s]

12 - Rise Time       [s]

13 - Pos Pulse Width [s]

14 - Neg Pulse Width [s]

15 - Dutycycle       [%]

reference

Variable Type       ViInt16

This control selects the reference channel for the measurement.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Edit) out of range.

BFFC0003  Parameter 3 (Measurement) out of range.

BFFC0004  Parameter 4 (Reference) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_setOptionDisplay

ViStatus mxOX2000_setOptionDisplay (ViSession instrumentHandle, ViBoolean grid,

ViInt16 color, ViBoolean screen,

ViInt16 menu, ViInt16 language);

Purpose

This function configures the display options of the instrument.

The following settings are available :

- Grid

- Color

- Screen

- Menu

- Language

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

grid

Variable Type       ViBoolean

This control sets the grid display.

Valid Values:

VI_OFF (0) - Off

VI_ON  (1) - On   (Default Value)

color

Variable Type       ViInt16

This control selects the color of the display.

Valid Values:

0 - Dark   (Default Value)

1 - Light

2 - White

screen

Variable Type       ViBoolean

This control selects the number of screens displayed.

Valid Values:

VI_OFF (0) - 1   (Default Value)

VI_ON  (1) - 4

menu

Variable Type       ViInt16

This control sets the menus to be displayed or not.

Valid Values:

0  - Off   (Default Value)

1  - Display

2  - Time base

3  - Channel 1

4  - Channel 2

5  - Channel 3

6  - Channel 4

7  - Trigger coupling

8  - Trigger mode/source

9  - Trigger select

10 - Cursors

11 - Time measurement

12 - Voltage measurement

13 - Functions

14 - Utility

15 - Files

16 - Port config

17 - Hardcopy

18 - Calibrate

19 - View setup

20 - Memory

language

Variable Type       ViInt16

This control selects the language used on the instrument.

Valid Values:

0 - French

1 - English  (Default Value)

2 - German

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Grid) out of range.

BFFC0003  Parameter 3 (Color) out of range.

BFFC0004  Parameter 4 (Screen) out of range.

BFFC0005  Parameter 5 (Menus) out of range.

BFFC0006  Parameter 6 (Language) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_setVideoTrigger

ViStatus mxOX2000_setVideoTrigger (ViSession instrumentHandle,

ViInt16 triggerSource, ViBoolean standard,

ViBoolean polarity, ViBoolean field,

ViInt16 line);

Purpose

This function configures the video trigger mode of the instrument.

The following settings are available :

- Trigger source

- Standard

- Polarity

- Field

- Line

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

triggerSource

Variable Type       ViInt16

This control selects the trigger source.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

standard

Variable Type       ViBoolean

This control selects the standard.

Valid Values:

VI_OFF (0) - 525

VI_ON  (1) - 625  (Default Value)

polarity

Variable Type       ViBoolean

This control selects the polarity.

Valid Values:

VI_OFF (0) - Pos (+)  (Default Value)

VI_ON  (1) - Neg (-)

field

Variable Type       ViBoolean

This control selects the field.

Valid Values:

VI_OFF (0) - Odd  (Default Value)

VI_ON  (1) - Even

line

Variable Type       ViInt16

This control selects the number of the line.

Valid Range:

Sandard      Field      Range

525          even       1 to 262

525          odd        1 to 263

625          even       314 to 625

625          odd        1 to 313

Default Value: 1

Note : This control value must be consistent with the field and

standard values. See the Digiscope 2000 Operating Instructions manual

for more details. If it is inconsistent an error is generated :

VI_ERROR_INCONSISTENT_PARAMETER6 (0xBFFC0906)

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Trigger Source) out of range.

BFFC0003  Parameter 3 (Standard) out of range.

BFFC0004  Parameter 4 (Polarity) out of range.

BFFC0005  Parameter 5 (Field) out of range.

BFFC0006  Parameter 6 (line) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFC0906  Parameter 6 (Line) is inconsistent.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_setWaveformDisplay

ViStatus mxOX2000_setWaveformDisplay (ViSession instrumentHandle, ViInt16 mode,

ViInt16 averagerate, ViInt16 persistence,

ViInt16 interpolation);

Purpose

This function configures the waveform display of the instrument.

The following settings are available :

- Mode

- Average rate

- Persistence

- Interpolation

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

mode

Variable Type       ViInt16

This control selects the display mode.

Valid Values:

0 - Normal  (Default Value)

1 - Dot join

2 - Persistence

3 - MinMax

4 - Envelope

averagerate

Variable Type       ViInt16

This control sets the average rate.

Valid Values:

0 - 1  (Default Value)

1 - 2

2 - 4

3 - 8

4 - 16

5 - 32

6 - 64

persistence

Variable Type       ViInt16

This control sets the persistence.

Valid Values:

0 - 200 ms  (Default Value)

1 - 500 ms

2 - 1 s

3 - 2 s

4 - 5 s

5 - 10 s

6 - inf

interpolation

Variable Type       ViInt16

This control selects the interpolation.

Valid Values:

0 - None  (Default Value)

1 - Linear

2 - Smooth

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Mode) out of range.

BFFC0003  Parameter 3 (Average rate) out of range.

BFFC0004  Parameter 4 (Persistence) out of range.

BFFC0005  Parameter 5 (Interpolation) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_verticalSetup

ViStatus mxOX2000_verticalSetup (ViSession instrumentHandle, ViInt16 input,

ViBoolean state, ViInt16 coupling,

ViInt16 probe, ViBoolean polarity,

ViInt16 voltDiv, ViReal64 offsetvolts,

ViBoolean bandwidth);

Purpose

This function configures the selected input channel of the instrument.

The following settings are available :

- Channel

- State

- Coupling

- Probe

- Channel inversion

- Range

- Offset

- Bandwidth limitation

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

input

Variable Type       ViInt16

This control selects the channel to configure.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

state

Variable Type       ViBoolean

This control selects the state of the specified channel.

Valid Values:

VI_ON  (1) - On  (Default Value)

VI_OFF (0) - Off

coupling

Variable Type       ViInt16

This control selects the coupling of the specified channel.

Valid Values:

0 - AC

1 - DC  (Default Value)

2 - GND

probe

Variable Type       ViInt16

This control selects the probe of the specified channel.

Valid Values:

1 - x 1  (Default Value)

2 - x 10

3 - x 100

polarity

Variable Type       ViBoolean

This control selects the polarity of the specified channel.

Valid Values:

VI_ON  (1) - Normal  (Default Value)

VI_OFF (0) - Inverted

voltDiv

Variable Type       ViInt16

This control sets vertical sense for the specified channel.

Valid Values:

0  - 2mV/div

1  - 5mV/div

2  - 10mV/div

3  - 20mV/div

4  - 50mV/div

5  - 100mV/div

6  - 200mV/div

7  - 500mV/div

8  - 1V/div

9  - 2V/div

10 - 5V/div

11 - 10V/div  (Default Value)

12 - 20V/div

13 - 50V/div

14 - 100V/div

15 - 200V/div

16 - 500V/div

17 - 1kV/div

Note : This control value must be consistent with the probe value.

See the Digiscope 2000 Operating Instructions manual for more

details. If it is inconsistent an error is generated :

VI_ERROR_INCONSISTENT_PARAMETER7 (0xBFFC0907)

offsetvolts

Variable Type       ViReal64

This control sets the signal offset of the specified channel.

Valid Ranges:

vertical sense       offset range

2mV/div              +/- 80mV

5mV/div              +/- 80mV

10mV/div             +/- 400mV

20mV/div             +/- 400mV

50mV/div             +/- 2V

100mV/div            +/- 2V

200mV/div            +/- 2.5V

500mV/div            +/- 12.5V

1V/div               +/- 12.5V

2V/div               +/- 64V

5V/div               +/- 64V

10V/div              +/- 64V

20V/div              +/- 640V

50V/div              +/- 640V

100V/div             +/- 640V

200V/div             +/- 6400V

500V/div             +/- 6400V

1kV/div              +/- 6400V

Default Value: 0.0

bandwidth

Variable Type       ViBoolean

This control selects the state of the specified channel.

Valid Values:

VI_ON  (1) - 150 MHz  (Default Value)

VI_OFF (0) - 20 MHz

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Input) out of range.

BFFC0003  Parameter 3 (State) out of range.

BFFC0004  Parameter 4 (Coupling) out of range.

BFFC0005  Parameter 5 (Probe) out of range.

BFFC0006  Parameter 6 (Polarity) out of range.

BFFC0007  Parameter 7 (Volts/div) out of range.

BFFC0008  Parameter 8 (Offset) out of range.

BFFC0803  Error scanning response.

BFFC0809  Parameter 9 (Bandwidth) out of range.

BFFC0900  Instrument specific error.

BFFC0907  Parameter 7 (Volts/div) is inconsistent.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_writeArrToFile

ViStatus mxOX2000_writeArrToFile (ViSession instrumentHandle,

ViReal64 waveformArray[],

ViInt16 numberofpoints, ViInt16 voltDiv,

ViReal64 offsetvolts, ViInt16 secDiv,

ViInt16 drives, ViString fileName,

ViBoolean overwriteFile);

Purpose

This function writes an array to a wave data file.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

waveformArray

Variable Type       ViReal64[]

This control specifies the waveform array to display.

Data are in volts.

numberofpoints

Variable Type       ViInt16

This control specifies the number of points to display.

voltDiv

Variable Type       ViInt16

This control sets the vertical sense for the data to be displayed.

Valid Values:

0  - 2mV/div

1  - 5mV/div

2  - 10mV/div

3  - 20mV/div

4  - 50mV/div

5  - 100mV/div

6  - 200mV/div

7  - 500mV/div

8  - 1V/div

9  - 2V/div

10 - 5V/div

11 - 10V/div  (Default Value)

12 - 20V/div

13 - 50V/div

14 - 100V/div

15 - 200V/div

16 - 500V/div

17 - 1kV/div

offsetvolts

Variable Type       ViReal64

This control sets the offset in volts for the data to be displayed.

Valid Ranges:

vertical sense       offset range

2mV/div              +/- 80mV

5mV/div              +/- 80mV

10mV/div             +/- 400mV

20mV/div             +/- 400mV

50mV/div             +/- 2V

100mV/div            +/- 2V

200mV/div            +/- 2.5V

500mV/div            +/- 12.5V

1V/div               +/- 12.5V

2V/div               +/- 64V

5V/div               +/- 64V

10V/div              +/- 64V

20V/div              +/- 640V

50V/div              +/- 640V

100V/div             +/- 640V

200V/div             +/- 6400V

500V/div             +/- 6400V

1kV/div              +/- 6400V

Default Value: 0.0

secDiv

Variable Type       ViInt16

This control sets the horizontal sense for the data to be displayed.

Valid Values:

0  - 2ns

1  - 5ns

2  - 10ns

3  - 20ns

4  - 50ns

5  - 100ns

6  - 200ns

7  - 500ns - (Default Value)

8  - 1s

9  - 2s

10 - 5s

11 - 10s

12 - 20s

13 - 50s

14 - 100s

15 - 200s

16 - 500s

17 - 1ms

18 - 2ms

19 - 5ms

20 - 10ms

21 - 20ms

22 - 50ms

23 - 100ms

24 - 200ms

25 - 500ms

26 - 1s

27 - 2s

28 - 5s

29 - 10s

30 - 20s

31 - 50s

drives

Variable Type       ViInt16

This control selects the target drive.

Valid Values:

0 - Computer  (Default Value)

1 - OX2000 Memory

2 - OX2000 PC-CARD

fileName

Variable Type       ViString

This function specifies the name of the wave data file.

Default Value: ""

Notes:

(1) When reading a file from Computer, precede all backslashes with

an additional backslash, e.g. if the file is in the \cvi\instr

directory, enter the filename as "\\cvi\\instr\\anyfile".

(2) File name must be 8-character long + ".TRC" extension.

If not an error is generated :0xBFFC0007

(3) If the target file already exists and overwrite is not selected

an error is generated : 0xBFFC090E

overwriteFile

Variable Type       ViBoolean

This function specifies whether or not to write over the specified

file if it already exists on disk.

Valid Range:

VI_OFF (0) - No  (Default Value)

VI_ON  (1) - Yes

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0003  Parameter 3 (Size Of Array) out of range.

BFFC0004  Parameter 4 (Volt/Div) out of range.

BFFC0005  Parameter 5 (Offset) out of range.

BFFC0006  Parameter 6 (timeBase) out of range.

BFFC0007  Parameter 7 (Drive) out of range.

BFFC0008  Parameter 8 (Filename) out of specifications.

BFFC0801  Error writing to file.

BFFC0803  Error scanning response.

BFFC0809  Parameter 9 (Overwrite File?) out of range.

BFFC0900  Instrument specific error.

BFFC090C  Channel is not active.

BFFC090E  File already exists.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_writeArrToMem

ViStatus mxOX2000_writeArrToMem (ViSession instrumentHandle,

ViReal64 waveformArray[],

ViInt16 numberofpoints, ViInt16 voltDiv,

ViReal64 offsetvolts, ViInt16 secDiv,

ViInt16 target);

Purpose

This function writes an array to the instrument display memory.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

waveformArray

Variable Type       ViReal64[]

This control specifies the waveform array to display.

Data are in volts.

numberofpoints

Variable Type       ViInt16

This control specifies the number of points to display.

voltDiv

Variable Type       ViInt16

This control sets the vertical sense for the data to be displayed.

Valid Values:

0  - 2mV/div

1  - 5mV/div

2  - 10mV/div

3  - 20mV/div

4  - 50mV/div

5  - 100mV/div

6  - 200mV/div

7  - 500mV/div

8  - 1V/div

9  - 2V/div

10 - 5V/div

11 - 10V/div  (Default Value)

12 - 20V/div

13 - 50V/div

14 - 100V/div

15 - 200V/div

16 - 500V/div

17 - 1kV/div

offsetvolts

Variable Type       ViReal64

This control sets the offset in volts for the data to be displayed.

Valid Ranges:

vertical sense       offset range

2mV/div              +/- 80mV

5mV/div              +/- 80mV

10mV/div             +/- 400mV

20mV/div             +/- 400mV

50mV/div             +/- 2V

100mV/div            +/- 2V

200mV/div            +/- 2.5V

500mV/div            +/- 12.5V

1V/div               +/- 12.5V

2V/div               +/- 64V

5V/div               +/- 64V

10V/div              +/- 64V

20V/div              +/- 640V

50V/div              +/- 640V

100V/div             +/- 640V

200V/div             +/- 6400V

500V/div             +/- 6400V

1kV/div              +/- 6400V

Default Value: 0.0

secDiv

Variable Type       ViInt16

This control sets the vertical sense for the data to be displayed.

Valid Values:

0  - 2ns

1  - 5ns

2  - 10ns

3  - 20ns

4  - 50ns

5  - 100ns

6  - 200ns

7  - 500ns - (Default Value)

8  - 1s

9  - 2s

10 - 5s

11 - 10s

12 - 20s

13 - 50s

14 - 100s

15 - 200s

16 - 500s

17 - 1ms

18 - 2ms

19 - 5ms

20 - 10ms

21 - 20ms

22 - 50ms

23 - 100ms

24 - 200ms

25 - 500ms

26 - 1s

27 - 2s

28 - 5s

29 - 10s

30 - 20s

31 - 50s

target

Variable Type       ViInt16

This control selects the target memory.

Valid Values:

1 - M1  (Default Value)

2 - M2

3 - M3

4 - M4

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0003  Parameter 3 (Size Of Array) out of range.

BFFC0004  Parameter 4 (Volt/Div) out of range.

BFFC0005  Parameter 5 (Offset) out of range.

BFFC0006  Parameter 6 (Target) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_writeFileToMem

ViStatus mxOX2000_writeFileToMem (ViSession instrumentHandle, ViInt16 drives,

ViString fileName, ViInt16 target);

Purpose

This function writes a wave data file to the instrument display memory.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

drives

Variable Type       ViInt16

This control selects the source drive.

Valid Values:

0 - Computer  (Default value)

1 - OX2000 Memory

2 - OX2000 PC-CARD

fileName

Variable Type       ViString

This control selects the target memory.

Default Value: ""

Notes:

(1) When reading a file from Computer, precede all backslashes with

an additional backslash, e.g. if the file is in the \cvi\instr

directory, enter the filename as "\\cvi\\instr\\anyfile".

(2) File name must be 8-character long + ".TRC" extension.

If not an error is generated :0xBFFC0003

(3) If file does not exist an error is generated : 0xBFFC090D

target

Variable Type       ViInt16

This control selects the target memory.

Valid Values:

1 - M1  (Default Value)

2 - M2

3 - M3

4 - M4

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Drive) out of range.

BFFC0003  Parameter 3 (Filename) out of specifications.

BFFC0004  Parameter 4 (Overwrite File?) out of range.

BFFC0800  Error opening file.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFC090D  File not found.

BFFC0910  Invalid file content.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mxOX2000_writeInstrData

ViStatus mxOX2000_writeInstrData (ViSession instrumentHandle,

ViString writeBuffer);

Purpose

This function writes commands and queries to the instrument to modify

parameters and query device settings.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

writeBuffer

Variable Type       ViString

The user can use this control to send common commands and queries to

the instrument. This control can also be used to write any valid

command to the instrument.

Default Value:  "*RST\n"

Notes:

(1) The command or query to be sent to the instrument may be a

literal enclosed in double quotes terminated by '\n' i.e. "*RST\n" or

may be contained in a variable of type string.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.